## Eye parameter
Analytical solution of personalized intraocular lens design for the correction of spherical aberration and coma of a pseudophakic eye
Bruno Chassagne1 and Lionel Canioni2
1ALPhANOV Centre Technologique Optique et Lasers, Institut d’Optique d’Aquitaine, 33400 Talence, France
2Université de Bordeaux, CNRS, CEA, CELIA, UMR 5107, F-33405 Talence, France
Eye values and IOL’s parameters used in the simulation
Surface | Radius (mm) | Thickness (mm) | Conic constant | 4th order aspheric term | n @ λ = 0.55 μm
Anterior Cornea | R1 | e1 | Q1 |  | n1
| 7.77 | 0.5 | -0.18 | NA | 1.376
Posterior Cornea | R2 | Postoperative aqueous depth e2 | Q2 |  | n2
| 6.4 | 4.1 | -0.6 | NA | 1.336
Anterior IOL | R3 | eIOL |  | e4 | nIOL
| 14.062 | 1.1 | 0 | -0.00062682 | 1.46
Posterior IOL | R4 | eIOL |  | e4 (-) | nIOL
| -9.4068 | 1.1 | 0 | 0.00062682 | 1.46
Vitreous | - | Vitrous chamber depth, e3 | - | - | n2
1.336
Retina R5 = | -12 |  | 0 | -
IOL’s parameters used in the simulation
k1 (rad) | 0.095993 | YIOL | 5.5043 | XIOL | -0.19835
h1 | 2.25 | PIOL (D) | 22 | k1(°) | 5.5
Optical model parameters |  |  |  | Operating instruction:
Corneal power, PC | 0.042251 | h2 | 1.9187 | The red values must be provided from eye measurement and targeted IOL: index, thickness, optical power of the IOL, anterior chamber depth e2.The expected input are displayed in red in the shaded zones. Dark orange values are provided from LBME eye model and could be also ajusted if necessary. The calculated intra ocular lens parameters are displayed in green
Anterior corneal focal length fCanterior | 28.4349 | h3 | 1.8228
Xc | 1.2932 | hmean | 1.8707
YC | 1 | u'1 = u2 | 0.071157
Lagrange invariant, H | 0.21598 | u’2 | 0.10196

## Corneal Aberrations
Numerical values for SA’s corneal coefficients
a1C | a2C | a3C | a4C | a5C | a6C | a7C | a8C | a9C | a10C
-585.4551 | -0.44158 | 64.9431 | 0.36555 | 1856.0171 | 7.2222 | -165.8943 | -1885.886 | 62.8928 | 676.3779
Numerical values for coma’s corneal coefficients
p1C | p2C | p3C | p4C | p5C | p6C
-16.2358 | 0.44158 | -0.18277 | 41.4736 | -4.3912 | -15.6133
Optical model parameters
PC Corneal power | 0.042251 | YC | 1
Xc | 1.2932
SIC | 0.0082038
SIIC | 0.0032972 | (SIIC)* | 0.0018869
| Prefactor | SA | conicity
SIC part | 0.00024163 | 39.5436 | -0.0013513
SIIC part | 0.00048798 | 6.6822 | 3.6376e-05

## IOL Aberration
Numerical values for SA’s IOL coefficients
a1IOL | a2IOL | a3IOL | a4IOL
0.0060009 | -0.0082978 | 0.0052052 | 24.3946
Numerical values for coma’s IOL coefficients
b1 | b2 | b3 | b4
0.00029 | 0.0017293 | -0.0013907 | -0.43449
A1 | A2 | A3 | A4 | A5 | A6 | A7 | A8 | A9 | A10
0.00013398 | 0.00048629 | 0.00034382 | 3.5397e-05 | 0.0057856 | 0.13304 | -0.00043439 | 0.021819 | 0.0017147 | -0.0016327
c
0.039678
Optical model parameters
4th order aspheric term e4 | -0.00062682
XIOL | -0.19835
SIIOL | -0.0082038
SIIIOL | -0.0019947

## Solution
Numerical solution for IOL parmeters
XIOL Opt | -0.19835
4th order aspheric term e4 Opt | -0.00062682
R3 | 14.062
R4 | -9.4068